# Unknown HMAC Co.

# Laboratory Equipment Manual

October 17, 2008

Revised 3/10/2006

# Unknown HMAC Co.

# Laboratory Equipment Manual

# Table of Contents

| Equipment | Section |
| --- | --- |
| Preface | 1 |
| Equipment Summary | 2 |
| Sieves | 3 |
| Shakers | 4 |
| Specific Gravity of Fine Aggregate Equipment | 5 |
| Tensile Molds | 6 |
| Compression Machines | 7 |
| Fine Aggregate Angularity Equipment | 8 |
| Gyratory Compactor | 9 |
| Gyratory Molds | 10 |
| General Thermometers | 11 |
| General Balances | 12 |
| General Ovens | 13 |
| Water Baths | 14 |
| Timers | 15 |
| Calipers | 16 |
| Vacuum Pumps | 17 |
| In-House Calibration Equipment | 18 |

Revised 3/10/2006

# SECTION 1:

# Preface;

# The purpose of the Laboratory Equipment Manual is to provide a method of verifying and documenting the suitability of laboratory equipment for testing purposes in accordance with the Louisiana Department of Transportation and Development Hot Mix Asphalt Plant Laboratory accreditation program.

# Equipment Verification / Manual Organization Procedure:

# The intent of this manual is to provide a working document for laboratory personnel or certified equipment technicians to document the finds of routine inspections of laboratory testing equipment.  Guidelines for completion and maintenance of this manual are as follows:

# The Laboratory Supervisor shall review the Laboratory Equipment Calibration / Verification Summary at the beginning of each month and coordinate the necessary laboratory equipment verification equipment checks due.

# All documents are to be completed in ink and signed by the certifying technician.

# Current equipment calibration / verification documents are to be stored in the Laboratory Equipment Manual and available to laboratory personnel.

# A filing system shall be developed with a folder for each section of this manual and the preceding verification documents stored for a period of no less than 3 years.

Revised 3/10/2006

# SECTION 2:

# Laboratory Equipment Calibration / Verification Summary

| Unknown HMAC Co. |
| --- |
| Equipment / Maintenance | Frequency | Verification Procedure | Date Performed | Date Due |
| Sieves (condition / opening size) | 12 Mo. | HMAC 1 |  |  |
| Shakers (sieving thoroughness) | 12 Mo. | HMAC 2 |  |  |
| Specific Gravity of Fine Aggregate Equipment | 6 Mo. | HMAC 3 |  |  |
| Tensile Equipment (dimensions, temp.) | 6 Mo. | HMAC 4 |  |  |
| Compression Machines (standardize) | 12 Mo. | HMAC 5 |  |  |
| Fine Aggregate Angularity Equipment | 12 Mo. | HMAC 6 |  |  |
| Gyratory Compactor (pressure, angle, freq.) | 6 Mo. | HMAC 7 |  |  |
| Gyratory Molds / Ram / Base (dimensions) | 12 Mo. | HMAC 8 |  |  |
| General Thermometers (calibrate) | 6 Mo. | HMAC 9 |  |  |
| General Balances (standardize) | 12 Mo. | HMAC 10 |  |  |
| General Ovens (standardize) | 6 Mo. | HMAC 11 |  |  |
| Water Bath (size, temp., suspension) | 6 Mo. | HMAC 12 |  |  |
| Timers (verify) | 12 Mo. | HMAC 13 |  |  |
| Calipers / Internal Bore Gauge (verify) | 12 Mo. | HMAC 14 |  |  |
| Vacuum Pump (verify) | 12 Mo. | HMAC 15 |  |  |

# SECTION 3

# Sieves

Revised 3/10/2006

# EQUIPMENT VERIFICATION

SIEVE OPENING SIZES

| Size | Maximum (mm) | Minimum (mm) |
| --- | --- | --- |
| 1 ½ | 39.5 | 36.4 |
| 1 ¼ | 33.2 | 30.5 |
| 1 | 26.4 | 24.2 |
| ¾ | 20.1 | 18.4 |
| 5/8 | 17.0 | 15.5 |
| ½ | 13.31 | 12.11 |
| 3/8 | 10.16 | 9.20 |
| 5/16 | 8.58 | 7.75 |
| ¼ | 6.78 | 6.10 |
| No. 4 | 5.14 | 4.60 |
| No. 8 | 2.600 | 2.280 |
| No. 10 | 2.215 | 1.930 |
| No. 12 | 1.890 | 1.640 |
| No. 16 | 1.330 | 1.135 |
| No. 20 | 0.970 | 0.815 |
| No. 30 | 0.695 | 0.575 |
| No. 40 | 0.502 | 0.406 |
| No. 50 | 0.363 | 0.286 |
| No. 60 | 0.306 | 0.238 |
| No. 100 | 0.192 | 0.142 |
| No. 140 | 0.141 | 0.100 |
| No. 200 | 0.103 | 0.070 |

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

SIEVES

DOTD TR 113

VERIFICATION PROCEDURE HMAC 1

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Calipers # ______________,   Microscope # _____________

| Sieve ID #: | Size: | Condition OK? |
| --- | --- | --- |
| Measurements: |

| Sieve ID #: | Size: | Condition OK? |
| --- | --- | --- |
| Measurements: |

| Sieve ID #: | Size: | Condition OK? |
| --- | --- | --- |
| Measurements: |

| Sieve ID #: | Size: | Condition OK? |
| --- | --- | --- |
| Measurements: |

# SECTION 4

# Mechanical Shakers

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

MECHANICAL SIEVE SHAKER

DOTD TR 113

VERIFICATION PROCEDURE HMAC 2

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Balance # ______________

| Sieve | Retained Mass | Hand Shaking, Material Passing, g | Percent Passing |
| --- | --- | --- | --- |

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 5

SPECIFIC GRAVITY AND ABSORPTION OF FINE AGGREGATE

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

SPECIFIC GRAVITY AND ABSORPTION OF FINE AGGREGATE

AASHTO T84

VERIFICATION PROCEDURE HMAC 3

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Calipers # ______________,   Balance # _____________, Ruler # ________________

| Tamper Mass, |  |
| --- | --- |
| Diameter, |  |  |

| Cone |
| --- |
| Top I.D., |  |  |
| Bottom I.D., |  |  |
| Height, |  |  |
| Wall Thickness, |  |  |

General Condition OK? _______________

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 6

Tensile Equipment

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

TENSILE EQUIPMENT

DOTD TR322

VERIFICATION PROCEDURE HMAC 4

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Calipers # ______________, Calibrated Thermometer # ________________

| Freezer Identification Number |  |
| --- | --- |

| Time Readings | Calibrated Thermometer Reading | Freezer Thermometer Reading | Difference ( + or - ) (0 ± 9°F?) |
| --- | --- | --- | --- |

Measurement of Loading Strips:

Width of Loading Strips =Upper ___________, ____________, Lower____________, ______________(0.75 in.)

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 7

Compression Machines

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

COMPRESSION MACHINES

DOTD TR 322

VERIFICATION PROCEDURE HMAC 5

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Calipers # ______________,   Balance # _____________

Documentation provided by the LA DOTD District Laboratory

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 8

Fine Aggregate Angularity Equipment

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

FINE AGGREGATE ANGULARITY

DOTD TR 121

VERIFICATION PROCEDURE HMAC 6

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Calipers # ______________,   Balance # _____________

| Funnel ID # |  |
| --- | --- |
| Opening Diameter, 90° (12.7 ± 0.6mm) |  |
| Opening Diameter, 180° (12.7 ± 0.6mm) |  |

Height from funnel to the top of the cylindrical measure = ____________ (115 ± 2mm)

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 9

Gyratory Compactor

Revised 9/22/2008

# EQUIPMENT VERIFICATION RECORD

GYRATORY COMPACTOR

AASHTO T312

CALIBRATION PROCEDURE HMAC 7

Calibration provided in-house by utilizing the manufacturers recommended procedure or performed by manufacturers certified technician.

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 10

Gyratory Compaction Molds

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

GYRATORY COMPACTOR MOLDS

AASHTO T312

CALIBRATION PROCEDURE HMAC 8

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency 12 Months

Verification Equipment Used:   Internal Bore Gauge # ______________,   Balance # _____________

| Mold ID |  |
| --- | --- |
| Mold Diameter, 149.90 to 150.00 |  |  |
| Mold Height, 250mm min. |  |  |
| Wall Thickness, 7.5mm min. |  |  |
| Ram Diameter, 149.50 to 149.75mm |  |  |
| Base Plate Diameter, 149.50 to 149.75mm |  |  |

| Mold ID |  |
| --- | --- |
| Mold Diameter, 149.90 to 150.00 |  |  |
| Mold Height, 250mm min. |  |  |
| Wall Thickness, 7.5mm min. |  |  |
| Ram Diameter, 149.50 to 149.75mm |  |  |
| Base Plate Diameter, 149.50 to 149.75mm |  |  |

| Mold ID |  |
| --- | --- |
| Mold Diameter, 149.90 to 150.00 |  |  |
| Mold Height, 250mm min. |  |  |
| Wall Thickness, 7.5mm min. |  |  |
| Ram Diameter, 149.50 to 149.75mm |  |  |
| Base Plate Diameter, 149.50 to 149.75mm |  |  |

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 11

Thermometers

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

# THERMOMETERS

# GENERAL

# VERIFICATION PROCEDURE HMAC 9

# Inspected By:__________________________________________________________ Date: ______/______/______

# Prev. Inspection Date: ______/______/______ Next Due Date: ______/______/______Verification frequency 6 months

Calibrated thermometer I.D. No.______________________________________________________________________

Test thermometer I.D. No.__________________________________________________________________________

Test thermometer use:______________________________________________________________________________

Test thermometer location:__________________________________________________________________________

Calibrated thermometer temp. reading:________________________________________________________________

Test thermometer temp. reading:______________________________________________________________________

Correction factor for thermometer:____________________________________________________________________

Thermometer tagged accordingly:_____________________________________________________________________

Calibrated thermometer I.D. No.______________________________________________________________________

Test thermometer I.D. No.__________________________________________________________________________

Test thermometer use:______________________________________________________________________________

Test thermometer location:__________________________________________________________________________

Calibrated thermometer temp. reading:________________________________________________________________

Test thermometer temp. reading:______________________________________________________________________

Correction factor for thermometer:____________________________________________________________________

Thermometer tagged accordingly:_____________________________________________________________________

Calibrated thermometer I.D. No.______________________________________________________________________

Test thermometer I.D. No.__________________________________________________________________________

Test thermometer use:______________________________________________________________________________

Test thermometer location:__________________________________________________________________________

Calibrated thermometer temp. reading:________________________________________________________________

Test thermometer temp. reading:______________________________________________________________________

Correction factor for thermometer:____________________________________________________________________

Thermometer tagged accordingly:_____________________________________________________________________

Calibrated thermometer I.D. No.______________________________________________________________________

Test thermometer I.D. No.__________________________________________________________________________

Test thermometer use:______________________________________________________________________________

Test thermometer location:__________________________________________________________________________

Calibrated thermometer temp. reading:________________________________________________________________

Test thermometer temp. reading:______________________________________________________________________

Correction factor for thermometer:____________________________________________________________________

Thermometer tagged accordingly:_____________________________________________________________________

Comments: __________________________________________________________________________________

# SECTION 12

Balances

Revised 9/22/2006

# EQUIPMENT VERIFICATION RECORD

BALANCES

GENERAL

CALIBRATION PROCEDURE HMAC 10

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Calipers # ______________,   Balance # _____________

# Calibration documentation provided by an independent scale service certified by the Department of Agriculture.

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 13

# Ovens

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

# OVEN TEMPERATURE SETTINGS

# VERIFICATION PROCEDURE HMAC 11

# Inspected By:____________________________________________________________ Date: _____/_____/_____

# Previous Verification Date: _____/_____/_____ Next Due Date: _____/_____/_____ Verification Frequency 6 months

# Calibrated Thermometer I.D. No.__________________________________________________________________

| Oven Identification Number |  |
| --- | --- |
| Oven Make |  |

| Time Readings | Calibrated Thermometer Reading | Oven Thermometer Reading | Difference ( + or - ) |
| --- | --- | --- | --- |

| Oven Identification Number |  |
| --- | --- |
| Oven Make |  |

| Time Readings | Calibrated Thermometer Reading | Oven Thermometer Reading | Difference ( + or - ) |
| --- | --- | --- | --- |

Comments: __________________________________________________________________________________

# SECTION 14

# Water Baths

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

WATER BATHS

GENERAL

VERIFICATION PROCEDURE HMAC 12

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Calipers # ______________,   Thermometer # _____________

TR304

| Water Bath ID |  |
| --- | --- |
| Temperature Readings (15min. intervals) 25 ± 1°C (77 ± 1.8°F)? |
| Overflow Outlet OK? | Suspension Apparatus OK? |

TR322

| Water Bath ID |  |
| --- | --- |
| Temperature Readings (15min. intervals) 25 ± 1°C (77 ± 1.8°F)? |

TR322

| Water Bath ID |  |
| --- | --- |
| Temperature Readings (15min. intervals) 60 ± 0.5°C (140 ± 1°F)? |

T85

| Water Bath ID |  |
| --- | --- |
| Temperature Readings (15min. intervals) 23 ± 1.7°C (73.4 ± 3°F)? |

Comments: __________________________________________________________________________________

# SECTION 15

# Timers

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

TIMERS

GENERAL

VERIFICATION PROCEDURE HMAC 13

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

# Certified Stopwatch I.D. Number..........................................................................___www.NIST.gov_______________

# CALIBRATION:

A = Reading on lab timer (sec)

B = Reading on certified timer (sec)

# Minimum Percent Accuracy = 0.05% when tested at least 20 minutes

| Laboratory Timer I.D. Number |  |
| --- | --- |
| Certified Timer Reading |  |
| Laboratory Timer Reading |  |
| Percent Accuracy |  |

| Laboratory Timer I.D. Number |  |
| --- | --- |
| Certified Timer Reading |  |
| Laboratory Timer Reading |  |
| Percent Accuracy |  |

| Laboratory Timer I.D. Number |  |
| --- | --- |
| Certified Timer Reading |  |
| Laboratory Timer Reading |  |
| Percent Accuracy |  |

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 16

Calipers

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

CALIAPERS

GENERAL

VERIFICATION PROCEDURE HMAC 14

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Gauge Block # ______________

Caliper Reading @ 50.00mm: _____________________

Caliper Reading @ 100.00mm: ____________________

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 17

Vacuum Pumps

Revised 3/10/2006

# EQUIPMENT VERIFICATION RECORD

VACUUM PUMP

GENERAL

VERIFICATION PROCEDURE HMAC 15

# Inspected By: _______________________________________________________Date: _______/_______/_______

# Previous Verification Date: _____/_____/____ Next Due Date: _____/_____/____Verification Frequency  12 Months

Verification Equipment Used: Pressure Gauge # ______________

Pump Verification:

# Reading on Vacuum Gauge: ______________

# Manometer Verification:

# Reading on Manometer: __________________

# Reading on Pressure Gauge: __________________

Comments: ______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

# SECTION 18

# In-House Calibration Equipment

# In-House Calibration Equipment

# Insert calibration equipment documentation in this section.

# (Suggested List of Items)

| Equipment | Approximate Cost |
| --- | --- |
| Digital Temperature Verification | 350.00 |
| Digital Bore Gauge | 2053.00 |
| Lenscope, Reticle | 107.00 |
| Sieve Opening Verification | 244.00 |
| Absolute Pressure Gauge, NIST Traceable | 465.00 |
| Calipers | 233.00 |
| End Measuring Rod, 50mm | 20.00 |
| End Measuring Rod, 125mm | 20.00 |
| Steel Rule | 10.00 |
| Thermometer Thermowell | Home Made |
| TOTAL | 3502.00 (note 1) |

Note 1:  Approximate Cost is list price for items.

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# SIEVES

# VERIFICATION PROCEDURE HMAC 1

# I.	EQUIPMENT

- Caliper readable to 0.01 mm (use for 4.75-mm sieve and coarser).

- Eye comparator with a 0.1-mm scale or a magnifier (for use with sieves finer than 4.75 mm).

# II.	PROCEDURE

A.	For sieves having openings equal to or greater than 4.75 mm, select and measure, using the calipers, the dimensions of at least 4 or 5 sieve openings in each sieve to ensure that the openings in the wire cloth conform to the requirements in Table 1 of AASHTO M 92 (ASTM E 11). Be sure to include, in the selection, any openings that appear distorted or unusual in size. Measure each of the openings as the distance between parallel wires measured at the center of each opening. Measure each opening in both the x (horizontal) and y (vertical) directions. Record the measurements for each of the selected openings. If a sieve has less than 5 full openings, measure all full openings.

B.	For sieves smaller than 4.75 mm, inspect the sieve cloth against a uniformly illuminated background. Use the eye comparator or magnifier to examine any suspicious areas of the cloth. If obvious deviations, such as weaving defects, creases, wrinkles, or excessive foreign matter in the cloth are found, the wire cloth is unacceptable.

C.	Inspect the general condition of the sieve. Check the frame and solder joints for cracks or holes. (Check for pin holes in the finer sieves.)

D.	Check for tightness of the wires on each individual sieve.

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# SHAKERS

# VERIFICATION PROCEDURE HMAC 2

# I.	EQUIPMENT

- Balance

# II.	PROCEDURE

A.	Perform a routine gradation using the normally allotted mechanical agitation time on the given shaker.

B.	Shake each sieve by hand for 1 minute and record the mass of material passing each sieve during the hand shaking procedure.

C.	Determine and record the mass retained on each sieve.

- Calculate and record the percent passing each sieve through manual agitation.

- Check for conformance.

Revised 3/10/2006

# EQUIPMENT VERIFICATION PROCEDURE

# APPARATUS FOR SPECIFIC GRAVITY

# AASHTO T-84

# VERIFICATION PROCEDURE HMAC 3

# I.	EQUIPMENT

A.	Calipers readable to 0.1 mm.

B.	Balance, 500 g capacity, readable to 0.1 g.

C.	Ruler readable to 1 mm.

# II.	PROCEDURE / CONE

A.	Measure the inside diameter at the top of the cone to the nearest 1 mm by taking two readings 90° apart using the ruler and record the results.

B.	Invert the cone and repeat step 1.

C.	Place the cone on a flat glass surface. Measure and record the depth of the cone.

D.	Using the calipers, measure the thickness of the cone to the nearest 0.1 mm by taking two readings 90° apart at the top of the cone and two readings 90° apart at the bottom of the cone and record the results.

III.	PROCEDURE / TAMPER

- Measure and record the diameter of the tamping face to the nearest 1 mm by taking two readings 90° apart using the ruler.

- Determine and record the mass of the tamper to the nearest 0.1 g.

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# TENSILE EQUIPMENT

# DOTD TR 322

# VERIFICATION PROCEDURE HMAC 4

# I.	EQUIPMENT

- Calipers readable to 0.1 mm.

B.	A calibrated thermometer graduated in 0.1C increments having a range which includes the temperature range to be checked.

C.	A brass thermometer well.

D.	A clothes pin to hold the thermometer

# II.	PROCEDURE

A.	Place the clothes pin on the thermometer.

B.	Place the thermometer inside the brass well with the clothes pin attached.  Position the thermometer brass well and clothespin on the shelf where the samples are normally dried.

C.	Take the first reading at least 1 hour after closing the oven (oven should remain undisturbed).  Record the time and temperature readings on both the calibrated thermometer and the oven thermometer.

D.	Take three consecutive readings no less than 1/2 hr. apart.  Determine if a temperature correction factor is necessary.

E.	Measure the tensile loading strips at two locations on each of the upper and lower strips.

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# COMPRESSION MACHINES

# DOTD TR 322

# VERIFICATION PROCEDURE HMAC 5

Verification provided by the LA DOTD District Laboratory

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# FINE AGGREGATE ANGULARITY EQUIPMENT

# DOTD TR 121

# VERIFICATION PROCEDURE HMAC 6

# I.	EQUIPMENT

- Calipers readable to 0.1 mm.

B.	Glass Plate, 60 x 60mm, 1/8” thick.

# II.	PROCEDURE

- Calibrate the volume of the unit weight measure cup according to DOTD TR 620.

- Measure the funnel opening sides on both the horizontal and vertical axis and record measure.

- Measure distance from bottom of funnel to the top of cylinder.

Revised 9/22/2006

# EQUIPMENT CALIBRATION PROCEDURE

# GYRATORY COMPACTOR

# AASHTO T 312

# VERIFICATION PROCEDURE HMAC 7

Method is dependant on individual facility:

1.	Calibration / Verification provided by manufacturers certified technician.

2.	Calibration can be performed in-house using the manufacturers recommended calibration procedure and equipment.

(Insert the manufacturer’s procedure within this section.)

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# GYRATORY COMPACTOR MOLDS

# AASHTO T 312

# VERIFICATION PROCEDURE HMAC 8

# I.	EQUIPMENT

A.	Calipers, capable of measuring the height of the molds and readable to 0.01 mm.

# II.	PROCEDURE

A.	Measure and record the inside diameter of the mold to the nearest 0.01 mm. Rotate the mold 90° (1/4 turn) and measure and record the inside diameter again.

B.	Measure and record the height of the mold to the nearest 0.01 mm. Rotate the mold 90° (1/4 turn) and measure and record the height again.

C.	Measure and record the thickness of the walls of the mold to the nearest 0.01 mm. Rotate the mold 90° (1/4 turn) and measure and record the wall thickness again.

D.	Measure and record the diameter of the ram to the nearest 0.01 mm. Rotate the mold 90° (1/4 turn) and measure and record the ram diameter again.

E.	Measure and record the diameter of the base plate to the nearest 0.01 mm. Rotate the mold 90° (1/4 turn) and measure and record the base plate diameter again.

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# THERMOMETERS

# GENERAL

# VERIFICATION PROCEDURE HMAC 9

# I.	EQUIPMENT

A.	Calibrated thermometer, same ASTM designation as test thermometer.

# II.	PROCEDURE

A.	Place the calibrated thermometer and the test thermometer at the same depth in the bath and as close together as possible.

B.	Allow at least fifteen minutes for the temperatures of the thermometers to stabilize.  Record the reading from both thermometers.

C.	If the readings are not identical on the thermometers calling for reference procedure E1, replace the test thermometer.

D.	On the baths where the reference procedure E77 is required, a corrected factor must be shown on a label on the bath or thermometer.

Revised 3/10/2006

# EQUIPMENT CALIBRATION PROCEDURE

# BALANCES

# GENERAL

# VERIFICATION PROCEDURE HMAC 10

# Calibration provided by a certified technician authorized by the Department of Agriculture.

Revised 3/10/2006

# EQUIPMENT VERIFICATION PROCEDURE

DRYING OVEN

GENERAL

VERIFICATION PROCEDURE HMAC 11

# I.	EQUIPMENT

A.	A calibrated thermometer graduated in 1.0C increments having a range which includes the temperature range to be checked.

B.	A brass thermometer well.

C.	A clothes pin to hold the thermometer.

# II.	PROCEDURE

A.	Place the clothes pin on the thermometer.

B.	Place the thermometer inside the brass well with the clothes pin attached.  Position the thermometer brass well and clothespin on the shelf where the samples are normally dried.

C.	Take the first reading at least 1 hour after closing the oven (oven should remain undisturbed).  Record the time and temperature readings on both the calibrated thermometer and the oven thermometer.

D.	Take three consecutive readings no less than 1/2 hr. apart.  Determine if a temperature correction factor is necessary.

F.	If a correction factor is found to be necessary; identify the installed oven thermometer and apply calibration factor when performing tests.

Revised 3/10/2006

# EQUIPMENT VERIFICATION PROCEDURE

WATER BATH

GENERAL

VERIFICATION PROCEDURE HMAC 12

# I.	EQUIPMENT

A.	A calibrated thermometer graduated in 1.0C increments having a range which includes the temperature range to be checked.

# II.	PROCEDURE

# A.	Inspect the water bath for appropriate capacity, overflow capabilities, and suspension apparatus.

# B.	Fasten the calibrated thermometer as close as practicable to the test thermometer and obtain and record the temperature of the water 4 times at 15 minute intervals.

Revised 3/10/2006

# EQUIPMENT VERIFICATION PROCEDURE

TIMERS

VERIFICATION PROCEDURE HMAC 13

# I.	EQUIPMENT

A.	Timer, readable to 0.1 sec., having a verified accuracy within the tolerance listed in test method.

“or”	www.nist.gov

# II.	PROCEDURE

A.	Start the timers simultaneously by pressing the "start" buttons at the same time.

“or”   Note time using www.nist.gov

B.	Run timers for a minimum of 15 min. and stop simultaneously.

C.	Record the two readings and calculate the percent accuracy.

Revised 3/10/2006

# EQUIPMENT VERIFICATION PROCEDURE

CALIPERS

GENERAL

VERIFICATION PROCEDURE HMAC 14

# I.	EQUIPMENT

- An end measuring rod, 50mm and 125mm.

II.	PROCEDURE

A.	Verify accuracy of calipers by measuring the 50mm x 125mm rods.

Revised 3/10/2006

# EQUIPMENT VERIFICATION PROCEDURE

VACUUM PUMP

GENERAL

VERIFICATION PROCEDURE HMAC 15

# I.	EQUIPMENT

A.	Certified Absolute pressure gauge.

B.	Water vapor trap

C.	Hoses, connectors, tools, misc.

# II.	PROCEDURE

A.	Connect the gauge to the system with the trap in-line between the system and gauge.

B.	Make sure all connections are air-tight.

- Open the number of lines normally used in testing and read and record the pressure indicated on the gauge.

D.	Repeat procedure utilizing the manometer.